'**************************************************

'Used to calibrate and adjust gauge mode PMMs.  Called by

'the master test macro "PMMCalSled_Master".

'******************************************************************************

Function FlukeCalSledGauge(iT, iL, iC, iP, cTest, cConfig)

Dim C1_Removed_Data()

Dim Reference_Data()

Dim C1_Corrected_Data()

Dim res(4,7)

C0 = 0

C1 = 0

zOffset = 0

'<currentteststep 1010 - record AF coefs.>

If cCOMPASS.CurrentTestStep = 1010 Then ' Test Definition data files were created for each DUT.

cDebug.LogStatus "Capture As Found Coeffs"

For i = 1 To cConfig.DUTPrs.Count

cDebug.LogStatus "DUT: " & i

' Read A/F C0 from PMM; write to calcoef1:

C0 = cConfig.DUTPrs(CInt(i)).IoSendCommand("OFFSET_SET?", False)

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient1 = C0

cDebug.LogStatus "Command OFFSET_SET?: " & C0

cDebug.LogStatus "CalibrationCoefficient1: " & cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient1

' Read A/F C1 from DUT; write to calcoef2

C1 = cConfig.DUTPrs(CInt(i)).IoSendCommand("GAIN_SET?", False)

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient2 = C1

cDebug.LogStatus "Command GAIN_SET?: " & C1

cDebug.LogStatus "CalibrationCoefficient2: " & cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient2

' Read A/F zOffset from PMM, write to calcoef3:

zOffset = cConfig.DUTPrs(CInt(i)).IoSendCommand("UCOEF_SET[0]?", False)

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient3 = zOffset

cDebug.LogStatus "Command UCOEF_SET[0]?: " & zOffset

cDebug.LogStatus "CalibrationCoefficient3: " & cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient3

calDate = cConfig.DUTPrs(CInt(i)).IoSendCommand("CAL_DATE?", False)

cCOMPASS.DataCollection(i).DUT.LastCalDate = calDate

cDebug.LogStatus "Command CAL_DATE?: " & calDate

cDebug.LogStatus "LastCalDate: " & cCOMPASS.DataCollection(i).DUT.LastCalDate

Next

cdebug.LogStatus "End of CurrentTestStep 1010"

'<currenteststep 1100 - zero the PMM>

ElseIf cCOMPASS.CurrentTestStep = 1100 And iC = 1 Then  ' New test pressure cycle and the 1st one

cDebug.LogStatus "Starting step 1100"

'DaqType mode = 0: manual Set device

DAQ = cCOMPASS.cConfig.SetPrs(1).RangeMain.GetParent.DaqType

cdebug.LogStatus "Data Acquisition Type for SetPrs is: " & DAQ

If DAQ <> 0 Then

cCOMPASS.StatusDisplay "Venting the pressure..."

cCOMPASS.cConfig.SetPrs(1).IoSetOutput 0,0,1

cDebug.LogStatus "Sent ioSetOutput 0,0,1 to vent"

TimeDelay 2

ElseIf DAQ = 0 Then

msg = msgbox("Vent the Pressure, Press OK when fully vented.",0+64+4096,"Manual Vent")

cCOMPASS.StatusDisplay "10s Dwell for Stabilization..."

TimeDelay 7

End If

For i = 1 To cConfig.DUTPrs.Count

'reset the Zoffset to "0.000", however this does not fully "zero" out the PMM as the C0 coef

'Is still being applied...

cCOMPASS.StatusDisplay "Removing the zero offset..."

TimeDelay 2

a = cCOMPASS.cConfig.DUTPrs(i).IoSendCommand("UCOEF_SET[0] 0",True)

TimeDelay 2

cDebug.LogStatus "sent cmd to zero Zoffset"

b = cCOMPASS.cConfig.DUTPrs(i).IoSendCommand("VAL?",False)

cDebug.LogStatus "VAL? return string is: " & b

'To do math on the pressure value it must be in a Data varient type.  Must first extract the

'number part from the string and then convert to a double precision data variant.

PMMValue = qextract(b,0,1,",")

cDebug.LogStatus "the PMMValue is: " & PMMValue

'convert the string to a data varient:

c = CDbl(PMMValue)

cDebug.LogStatus "the data type value for PPMValue is: " & c

newOffset = c * -1

cDebug.LogStatus "the newOffset value is: " & newOffset

'finish the zeroing process...

cmd = "UCOEF_SET[0] " & newOffset

e = cCOMPASS.cConfig.DUTPrs(i).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "offset command sent to the PMM: " & cmd

'verify new zoffset correctly received by PMM, save it to DUTRaw3:

zOffset = cConfig.DUTPrs(i).IoSendCommand("UCOEF_SET[0]?", False)

cCOMPASS.DataCollection(i).DUTRaw3.CalibrationCoefficient3 = zOffset

Next

cDebug.LogStatus "End of CurrentTestStep 1100"

'Continue with pressure cycle 1 (as-found data)

'<currenteststep 1150 - pressure cycle 1 is done, run polyfit>

ElseIf cCOMPASS.CurrentTestStep = 1150 And iC = 1 Then

cCOMPASS.StatusDisplay "Calculating Adjustment..."

cDebug.LogStatus "First Cycle is complete. Starting CurrentTestStep 1150 - run PolyFit"

cDebug.LogStatus "#Data files: " & cCOMPASS.DataCollection.Count

For i = 1 To cConfig.DUTPrs.Count

pressurePoints = cCOMPASS.DataCollection(i).NumberofPressurePoints

cDebug.LogStatus "DUT: " & i & " - Pressure Points: " & pressurePoints

Redim Reference_Data(pressurePoints-1)

Redim C1_Removed_Data(pressurePoints-1)

Redim C1_Corrected_Data(pressurePoints-1)

'populate PMM A/F coefficients from memory:

C0 = cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient1

C1 = cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient2

zOffset = cCOMPASS.DataCollection(i).DUTRaw3.CalibrationCoefficient3

cDebug.LogStatus "C0: " & C0

cDebug.LogStatus "C1: " & C1

cDebug.LogStatus "zOffset: " & zOffset

dutUnit = cConfig.DUTPrs(i).RangeMain.UnitFinal

cDebug.LogStatus "DUT UnitFinal: " & dutUnit

cDebug.LogStatus "DUT UnitFinalText: " & cConfig.DUTPrs(i).RangeMain.UnitFinalText

refUnit = cConfig.RefPrs(i).RangeMain.UnitFinal

cDebug.LogStatus "Ref UnitFinal: " & refUnit

cDebug.LogStatus "Ref UnitFinal: " & cConfig.RefPrs(i).RangeMain.UnitFinalText

' Iterate through each pressure point

For j = 1 To pressurePoints

ix = CInt(j)-1

cDebug.LogStatus "FlukeCalSledCalibrationAdjust DUTPressurePoint(" & j & ")"

DUTPressure = cCOMPASS.DataCollection(i).DataPointRef(1, 1, 1, CInt(j)).DUTPressure

cDebug.LogStatus "FlukeCalSledCalibrationAdjust DUTPressurePoint: " & DUTPressure

DUTPsiPres = cCOMPASS.UnitConversion(CDbl(DUTPressure), 9, CInt(dutUnit), 0)

cDebug.LogStatus "FlukeCalSledCalibrationAdjust DUTPressurePoint (psi): " & DUTPsiPres

refPres = cCOMPASS.DataCollection(i).DataPointRef(1, 1, 1, CInt(j)).RefPressure

cDebug.LogStatus "FlukeCalSledCalibrationAdjust ReferencePressurePoint: " & refPres

Reference_Data(ix) = cCOMPASS.UnitConversion(CDbl(refPres), 9, CInt(refUnit), 0)

cDebug.LogStatus "FlukeCalSledCalibrationAdjust ReferencePressurePoint (psi): " & Reference_Data(ix)

zOff_Removed = CDbl(DUTPsiPres) - CDbl(zOffset)

cDebug.LogStatus "FlukeCalSledCalibrationAdjust zOff_Removed: " & zOff_Removed

C0_Removed = zOff_Removed - C0

cDebug.LogStatus "FlukeCalSledCalibrationAdjust C0_Removed: " & C0_Removed

C1_Removed_Data(ix) = C0_Removed / C1

cDebug.LogStatus "FlukeCalSledCalibrationAdjust C1_Removed_Data: " & C1_Removed_Data(ix)

Next

Call Poly_Fit(Reference_Data, C1_Removed_Data, res, 1)

slope= res(0,1)

cDebug.LogStatus "FlukeCalSledCalibrationAdjust slope: " & slope

If slope = 0 Then

new_C1 = 0

Else

new_C1 = 1 / slope

End If

cDebug.LogStatus "FlukeCalSledCalibrationAdjust new_C1: " & new_C1

'Calculate the C1_corrected for each point:

For j = 0 To pressurePoints - 1

C1_Corrected_Data(CInt(j)) = C1_Removed_Data(CInt(j)) * new_C1

cDebug.LogStatus "FlukeCalSledCalibrationAdjust C1_Corrected_Data(" & j & "): " & C1_Corrected_Data(CInt(j))

Next

Call Poly_Fit(Reference_Data, C1_Corrected_Data, res, 1)

new_C0 = 0 - res(0,0)

cDebug.LogStatus "FlukeCalSledCalibrationAdjust new_C0: " & new_C0

'Save new coefs to memory AND write to the PMM:

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient4 = new_C0

cmd = "OFFSET_SET " & new_C0

Call cConfig.DUTPrs(CInt(i)).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "CalibrationCoefficient4: " & cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient4

cDebug.LogStatus "Sent command: " & cmd

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient5 = new_C1

cmd = "GAIN_SET " & new_C1

Call cConfig.DUTPrs(CInt(i)).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "CalibrationCoefficient5: " & cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient5

cDebug.LogStatus "Sent command: " & cmd

'Determine new date of calibration, save to memory AND write to PMM:

new_date = FlukeCalSled_Date

cCOMPASS.DataCollection(i).DUT.CalDueDate = CDate(new_date)

cmd = "CAL_DATE " & new_date

Call cConfig.DUTPrs(CInt(i)).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "CalDueDate: " & cCOMPASS.DataCollection(i).DUT.CalDueDate

cDebug.LogStatus "Sent command: " & cmd

'Reset Zoffset to zero:

cmd = "UCOEF_SET[0] 0"

Call cConfig.DUTPrs(CInt(i)).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "Sent command: " & cmd

Next

Set obj = Nothing

cdebug.LogStatus "Finished with CurrentTestStep 1150"

'<currentteststep 2000 - test is complete, AL data has been collected. Save the coefs>

' Test complete. The post test options have not displayed and the user notes have not been entered.

ElseIf cCOMPASS.CurrentTestStep = 2000 Then

cDebug.LogStatus "Test complete.  CurrentTestStep 2000.  Get final zOffset"

'Capture final zOffset - should be zero

For i = 1 To cConfig.DUTPrs.Count

cDebug.LogStatus "DUT: " & i

' Read zOffset from DUT

zOffset = cConfig.DUTPrs(CInt(i)).IoSendCommand("UCOEF_SET[0]?", False)

cDebug.LogStatus "Command UCOEF_SET[0]?: " & zOffset

' Write the final zOffset value

cCOMPASS.DataCollection(i).DUT.CalibrationCoefficient6 = zOffset

Next

'OPTIONAL GUARD BANDING FEATURE:  This block of code can be commented out if guardbanding evaluation is not desired.

TOLFACT = .50   '50% gaurdband.  User definable.

For i = 1 To cCOMPASS.DataCollection(1).NumberofPressurePoints

For j = 1 To cConfig.DUTPrs.Count

Ref = cCOMPASS.DataCollection(j).DataPointRef(1,1,2,CInt(i)).RefPressure

DUT = cCOMPASS.DataCollection(j).DataPointRef(1,1,2,CInt(i)).DUTPressure

Tol  = cCOMPASS.DataCollection(j).DataPointRef(1,1,2,CInt(i)).Tolerance

ST = cCOMPASS.DataCollection(j).DataPointRef(1,1,2,CInt(i)).Status

If abs(DUT-Ref) > Tol * TOLFACT Then

FAIL = True

cCOMPASS.DataCollection(j).DataPointRef(1,1,2,CInt(i)).Status = ST & "T"

'Generate Failure message for user

msgbox "The As Left results did not fall inside the guardbanded tolerance. Click OK to continue."

Exit For

End If

Next

If fail = True Then Exit For

'END OF GUARD BANDING CODE.

Next

msg = "The test is complete.  The coefficients are active in module memory but "

msg = msg & "have not been written to permanent storage. Would you like to activate "

msg = msg & "the calibration?"

If msgbox(msg,vbquestion + vbYesNo + vbSystemModal,"Activate Calibration") = vbNO Then

msg = "Recycling the power of the module will "

msg = msg & "remove the coefficients from the memory and reset the module."

msgbox msg,vbSystemModal,"Activate Calibration"

Exit Function

Else

For i = 1 To cConfig.DUTPrs.Count

cmd = "CAL_STORE"

TimeDelay 2

Call cConfig.DUTPrs(CInt(i)).IoSendCommand(CStr(cmd), True)

cDebug.LogStatus "Sent command: " & cmd

Next

End If

End If

End Function