If you’re using a COM port other than COM1, then change the default COM port here in the Interface macro named “Fluke700P_GetPressure”. Save the changes.

o

If this line is not already commented out (by adding an apostrophe at the beginning of it), then comment it out. Make sure it turns green. Save the changes.

If you are using the example or sample 750P module DUT setup in COMPASS, then uncheck this box in the [Database], <Setup> menu. If you don’t uncheck it, then any changes that you make to the macros related to it will revert automatically!

Interface macro “Fluke700P_GetPressure” edited 2018-Dec-04 by Kyle Clark with the above changes.

'****************************************

'This Function must read and return the output of a remote instrument.

'The output must be in the Raw Output Unit defined by the Range class.

'

'cCalc     :The calculation class of the device.

'cParent   :The parent device class.

'fTarget1  :Target output for controller.

'bChnlSt   :True/False to activate or de-active a valve driver.

'cConfig   :Configuration class that holds all active calculation classes.

'

'The value is returned by setting the function name = 'to the calculated value..

'

'For example: Interface33489 =  val(mid(rawReply,5))

'****************************************

Function Fluke700P_GetPressure(cCalc, cParent, fTarget1, bChnlSt, cConfig)

On Error Resume Next

If CheckFluke700PObj(cCalc) Then

prs = cCalc.Obj1.GetPressureP()

'cDebug.LogStatus "700P Pressure: " & prs

'cDebug.LogStatus "Error Code:" & cCalc.Obj1.ErrorCode

dUnit = cCalc.RangeMain.UnitFinal

newPrs = cCOMPASS.UnitConversion(CDbl(prs), CInt(dUnit), 9, 0)

zeroPrs = cCalc.GetParamData(0)

Fluke700P_GetPressure = newPrs - zeroPrs

End If

End Function

'''''''''''''''''''''''''''''''''''''''''''''''''''

'Make sure that the Fluke700P object is set.

'This Function Is Not called directly

'by COMPASS.  It is used by the other Fluke700P

'related functions.

'''''''''''''''''''''''''''''''''''''''''''''''''''

Function CheckFluke700PObj(cCalc)

On Error Resume Next

defaultPort = 1	'If you're not using RS-232 COM1 then change to defaultPort = n where n Is the default COM port you want To use

'cDebug.LogStatus "Default Port:" & defaultPort

If cCOMPASS.SystemAbort Then

'If cCalc.Obj1 Is Nothing Then Exit Function

If cCalc.GetParamData(5) = 0 Then

temp = cCalc.Obj1.Errordescription

cCOMPASS.StatusDisplay "Shutting down 700 PCK interface..."

cdebug.LogStatus "Shutting down 700 PCK interface..."

cCalc.Obj1.Disconnect

cCOMPASS.SystemAbort = False

cCOMPASS.TimeDelay 5

cCOMPASS.SystemAbort = True

Set cCalc.Obj1 = Nothing

cDebug.LogStatus "PCK Object Error Description: " & temp

cDebug.LogStatus "Disconnected Fluke Interface. Err:" & Err

cCalc.SetParamData 5,1

End If

Exit Function

End If

If cCOMPASS.SystemAbort Then Exit Function

If cCalc.Obj1 Is Nothing And cCalc.GetParamData(1) = 0 Then

'Check to see if this is DUT #1

PortInfo = GetVal(cCOMPASS.cConfig.DUTPrs(1).RangeMain.GetParent.ParamID)

cDebug.LogStatus "*******Fluke Interface INIT********"

cDebug.LogStatus "Model: " & cCalc.RangeMain.GetParent.Model

cDebug.LogStatus "SN: " & cCalc.RangeMain.GetParent.SN

cDebug.LogStatus "COM PortInfo from macro:" & defaultPort

cDebug.LogStatus "COM PortInfo from Parameter ID:" & PortInfo

If PortInfo > 0 And PortInfo < 255 Then

Port =  PortInfo

cDebug.LogStatus "Port, after check for 255 > ParamID > 0 :" & Port

Else

Port = defaultPort

cDebug.LogStatus "Port, if no Parameter ID (from defaultPort):" & Port

End If

K1 = cConfig.DUTPrs(1).RangeMain.Key

P1 = cConfig.DUTPrs(1).RangeMain.GetParent.ProfileID

P2 = cCalc.RangeMain.GetParent.ProfileID

If (cCalc.RangeMain.Key = K1 And P2 = P1) Or cConfig.DUTPrs.Count = 1 Then

'First DUT

DUTID = 1

'Port =   PortInfo    '(Don't do this. It sets Port back to PortInfo from DUT #1 ParamID. If DUT #1 ParamID is blank this will be wrong.)

cCOMPASS.StatusDisplay "First DUT. Default COM port in macro is " & PortInfo & ". COM" & Port & " will be used for DUT#1"

cDebug.LogStatus "First DUT," & " SN " & cCalc.RangeMain.GetParent.SN & " on COM " & Port

Else

If cConfig.DUTPrs(1).Obj1 Is Nothing Then

cDebug.LogStatus "Wait for First DUT to init"

cDebug.LogStatus "Init Exit"

Exit Function

End If

For i = 2 To cConfig.DUTPrs.Count

If cCalc.RangeMain.Key = cConfig.DUTPrs(i).RangeMain.Key Then

PortInfo = cCOMPASS.cConfig.DUTPrs(i).RangeMain.GetParent.ParamID

Port = PortInfo 'This is correct. COM port for more than 1 DUT has to come from ParamID

DUTID = i

cCOMPASS.StatusDisplay "DUT#" & DUTID & ". ParamID value is " & Port & ". COM" & Port & " will be used for DUT# " & DUTID

cDebug.LogStatus " DUT # " & DUTID & ". ParamID = " & PortInfo & ", COM Port = " & Port

Exit For

End If

Next

End If

'Save the FlukeINterface object

'to the Calculation class of the device.

cCOMPASS.StatusDisplay "Updating Fluke Interface"

cDebug.LogStatus "Initializing Fluke Interface. Err:" & Err

Set cObj = GetObject("", "Fluke.Calibration.PCKsupport")

'****Specify Port to Communicate*****

Model = cCalc.RangeMain.GetParent.Model

SN = cCalc.RangeMain.GetParent.SN

cCOMPASS.StatusDisplay "Model:" & Model & "(SN" & SN & ") Initializing " & Port

'No property by the name ChangeComPort in the support library

cDebug.LogStatus "Setup to use " & Port & ". Err:" & Err

TimeDelay 3  'It takes a few seconds to update the  port

cCOMPASS.StatusDisplay "Connecting Fluke Interface"

cObj.Connect CInt(Port), CInt(rtv)

cDebug.LogStatus "Fluke 700 Connection"

cDebug.LogStatus "RS232 Port:" & Port

cDebug.LogStatus "Error Code:" & cObj.ErrorCode

cDebug.LogStatus "Error Description:" & cObj.ErrorDescription

If cObj.ErrorCode <> 0 Then  'len(temp)> 5 Then

cCOMPASS.SystemAbort = True

cCOMPASS.StatusDisplay "Connection failed on Port " & Port

cDebug.LogStatus "Connection Failed on Port " & Port

msg = "Connection to Fluke module failed on COM Port " & Port

msg = msg & ".  Please verify that the port is "

msg = msg & "valid and that the sensor is properly connected.  "

msg = msg & "Enter the proper COM port in the Parameter ID field "

msg = msg & "To specify a port other than " & Port & "."

msg = msg & vbcrlf & "ERR:" & vbcrlf & temp

MsgBox msg, vbCritical, "Error"

cCalc.SetParamData 1, 1 'Flag that device failed

cCOMPASS.SystemAbort = True

Exit Function

End If

Set cCalc.Obj1 = cObj

'Update Serial Number of device.

temp = cCalc.Obj1.SerialNumber

If Len(temp) > 2 Then

cCalc.RangeMain.GetParent.SN = temp

cDebug.LogStatus "Set SN to :" & temp

End If

'Update Model Info of device.

temp = cCalc.Obj1.Model

If Len(temp) > 2 Then

cCalc.RangeMain.GetParent.Model = temp

cDebug.LogStatus "Set Model to :" & temp

End If

'bottom sale and full scale

DUTMin = cCalc.RangeMain.MinFinal

DUTMax = cCalc.RangeMain.MaxFinal

cDebug.LogStatus "DUT Min = " & DUTMin & ",DUT Max = " & DUTMax

btmScale = cCalc.Obj1.BottomScale

fullScale = cCalc.Obj1.FullScale

cDebug.LogStatus "BottomScale = " & btmScale & ",FullScale = " & fullScale

'If data files were already created, add the Model and SN info the files

If (cCOMPASS.COMPASSRunState And 2 ^ 6) <> 0 Then

cCOMPASS.DataCollection(DUTID).DUT.Model = cCalc.RangeMain.GetParent.Model

cCOMPASS.DataCollection(DUTID).DUT.SerialNumber = cCalc.RangeMain.GetParent.SN

End If

cDebug.LogStatus "Connection active"

cCalc.SetParamData 1, 2 'Flag that the device is set...

ElseIf cCalc.GetParamData(1) = 1 Then

'Tried and failed

Exit Function

End If

CheckFluke700PObj = True

End Function