# Calibration Results

# Applied	Displayed

# psi	psi

# ----------	------------

# 0.0000	0.0000

# 1.0000	1.0000

# 2.0000	2.0000

# 3.0000	3.0001

# 4.0000	4.0001

# 5.0000	5.0000

# 6.0000	6.0000

# 7.0000	6.9999

# 8.0000	7.9999

# 9.0000	9.0000

# 10.0000	10.0000

# Statement of Accuracy

The standards used to calibrate this instrument are are maintained according to Corporate Quality Policy.  All measurements were conducted in a controlled environment with instrumentation adequate for assurance that the reported values are reliable.  The calibration process satisfies the requirements of the MILSTD45662A.

The horizontal plane of reference for pressure measurement was selected at the pressure reference mark on the front panel. The two sigma rss uncertainty in the pressure generated to calibrate this instrument is estimated to be less than or equal to 23 ppm of applied pressure.  The reported values are traceable to the National Institute of Standards and Technology through test numbers none and none.

Calibrated by: James J. Technician	Signature: _____________________